
 




Wide Range Frequency Response Compensation Using DSP

Laszlo Hars

Panasonic Information and Networking Technologies Laboratory

E-mail: Laszlo.Hars@Research.Panasonic.com
Abstract: When an instrument works over many narrow band frequency channels at non-flat hardware frequency response, simple FIR filters can be applied for compensation. A very fast method is proposed for run time design these filters based on calibration data.
Introduction

In modern telecommunication systems there are cases when an instrument has to work over hundreds of narrow band frequency channels. It is difficult and expensive to build hardware with flat frequency response all over them. 

Filter Table

When a particular center frequency for a channel is selected a simple FIR filter can be applied for compensation. The frequency response of this filter depends on the center frequency, which would require storing thousands of sets of the filter coefficient along with a table telling us which filter is needed where. The filters have to be determined at calibration time. If the center frequency can be chosen freely, this pre-calculated filter approach is impractical, requiring huge storage space and very long calibration time.

Run-time Filter Design

A better way is to measure the amplitude characteristics of the equipment on a sufficiently dense frequency grid when the instrument gets calibrated. Only this table has to be stored. Using interpolation (linear, cubic or spline – depending the smoothness of the frequency response curve) the required compensation filter-response can be determined with a handful of arithmetic operations at run time. The compensation filter can be very short if the sampling rate and intermediate frequency are chosen appropriately. We developed a closed form expression for the filter coefficients, which allows us to compute them with another handful operations. All together the calculations are so fast, that even arbitrary frequency hopping can be implemented with slow, low power DSPs.

Calibration Tables

The input signal has to be attenuated and/or amplified. These circuits are not perfect either. In theory, we need a frequency response table for each possible attenuation. In practice, however, the frequency-response changes only at smaller attenuation values, higher values provide good de-coupling between otherwise interfering circuit parts. These mean, that around 10 compensation tables are often enough for even high precision measurements. The affects of different attenuation devices are cumulative at proper design (at least at higher attenuation values), what further reduces the number of necessary tables.

Temperature compensation can be incorporated, too. The measured ambient or internal temperature represents another dimension for the family of tables.

FIR vs. IIR Filters

The IIR filters have more complicated formulae for their gain response than the FIR filters do, therefore the online calculation of the coefficients are longer. We normally also need constant group delay in the pass band, which is more difficult to achieve with IIR filters.

The filters must be stable, that is the roots of the denominator of the transfer function of IIR filters must all lie inside the complex unit circle. It is another difficulty to be dealt with.

On the other hand, the same long IIR filters can somewhat better approximate a given gain curve. If this curve has very sharp peaks, notches or edges, IIR filters are needed. Also, the FIR filters have large group delays. If the group delay must be small or even negative, IIR filters have to be used.

The little better approximation of the desired amplitude curve by an IIR filter can be balanced by applying a longer FIR filter, whose coefficient calculations are simpler. The filtering take about the same time because of the faster FIR filter-code.

Therefore FIR filters have been chosen. They must be short, having less than 10 coefficients, otherwise the calculation of the coefficients gets complicated.

Linear Phase FIR Filters

The frequency response function of a FIR filter with coefficient sequence 
[image: image1.wmf]c

c

c

N

0

1

1

,

,

K

-

 is given by


[image: image2.wmf]H

c

e

k

i

k

k

N

w

w

b

g

=

×

-

×

=

-

å

0

1

.

If the filter has linear phase response (constant group delay), the coefficient sequence must be symmetric or anti-symmetric. We want a filter with relative flat amplitude response, that is, close to unity gain everywhere, also at DC. If the coefficient sequence is anti-symmetric (
[image: image3.wmf]c

c

k

N

k

=

-

-

-

1

), the DC gain is 0. Therefore we need a symmetric coefficient sequence.

The filter delay is (N–1)/2. It is easier to compensate or handle an integer number of samples delay, what we have if N is odd. In this case the response function becomes


[image: image4.wmf]H

e

c

c

k

N

i

k

N

k

N

N

N

w

w

w

b

g

d

i

e

j

=

+

-

F

H

G

G

G

I

K

J

J

J

-

-

×

-

=

+

-

-

å

1

2

1

2

1

2

1

1

2

2

cos


The factor 
[image: image5.wmf]e

N

i

-

-

1

2

w

 represents the delay having a constant magnitude of 1, and the real factor in the brackets gives the (signed) amplitude response.

The length of the filter can be chosen to be 7. It has 4 free coefficients. We need the gain be a given value at 3 different frequencies and one degree of freedom remains to enforce a smooth amplitude response curve. Let the coefficient sequence be 
[image: image6.wmf]c

b

a

d

a

b

c

,

,

,

,

,

,

. The amplitude response of the corresponding FIR filter is


[image: image7.wmf]A

d

a

b

c

w

w

w

w

b

g

b

g

b

g

b

g

=

+

+

+

2

2

2

2

3

cos

cos

cos


Sampling and Signal Frequency

Usually, if some signal conversion is performed before the amplitude response correction we can simplify the design and save processing time. We need mixing and decimation to reduce the sampling rate 
[image: image8.wmf]f

Samp

 to- or a little above the double of the signal bandwidth. This is the minimum, which preserves all the information of the original signal (Nyquist theorem). It usually involves a band-pass filter step, too, removing those disturbing signals, which would alias to the useful frequency band. This frequency band is best located around the center frequency 
[image: image9.wmf]f

0

, 
[image: image10.wmf]f

f

f

Nyq

Samp

0

1

2

1

4

=

=

 (or at 
[image: image11.wmf]3

2

f

Nyq

), such that no signal component aliases back into the useful frequency band at another location. (On the normalized scale where the Nyquist frequency is 1, 
[image: image12.wmf]w

p

=

×

f

, 
[image: image13.wmf]f

0

1

2

=

.)

Filter Design

Applying the compensation filter we do not want to change the amplitude at 
[image: image14.wmf]f

0

, the center of the frequency band of the signal. (If necessary, we adjust the gain outside of the filter.) For the frequency response compensation we specify the gains 
[image: image15.wmf]g

1

 and 
[image: image16.wmf]g

2

 of the filter at 2 other frequencies, say 
[image: image17.wmf]f

1

1

4

=

 and 
[image: image18.wmf]f

2

3

4

=

 at both sides of 
[image: image19.wmf]f

0

. (They can be chosen closer or further away from 
[image: image20.wmf]f

0

, according to the need to have more accurate compensation close to the center frequency or less accurate compensation over the whole band.) These represent 3 constraints for the filter response, which is a function of 4 free coefficients.

A very important additional requirement is, that the filter must not have large ripple, i.e. its amplitude response must be smooth. This can be guaranteed, if we require the slope of the response curve at 
[image: image21.wmf]f

0

 be the same as that of the secant line between 
[image: image22.wmf]f

1

 and 
[image: image23.wmf]f

2

.

These requirements give us a linear system of 4 equations for the 4 unknown coefficients of the filter.


[image: image24.wmf]1

2

2

2

2

2

2

2

2

2

2

3

2

3

2

1

4

2

3

4

2

3

4

3

2

9

4

1

2

2

3

2

=

+

+

+

=

+

+

+

=

+

+

+

-

=

+

+

d

a

b

c

g

d

a

b

c

g

d

a

b

c

g

g

a

b

c

cos

cos

cos

cos

cos

cos

cos

cos

cos

sin

sin

sin

p

p

p

p

p

p

p

p

p

p

p

p

p

p

p

d

i

b

g

d

i

d

i

d

i

d

i

d

i

d

i

d

i

d

i

b

g

d

i


With the real values of the trigonometric functions:


[image: image25.wmf]1

2

2

2

2

2

3

1

2

1

2

=

-

=

+

-

=

-

+

-

=

-

d

b

g

d

a

c

g

d

a

c

g

g

a

c

p

p


we get a simple solution:


[image: image26.wmf]a

g

g

b

g

g

c

g

g

d

g

g

=

-

×

-

F

H

G

I

K

J

=

+

×

-

=

-

×

-

F

H

G

I

K

J

=

+

×

1

2

1

2

1

2

1

2

3

2

8

1

2

1

4

1

2

2

8

1

2

1

2

b

g

b

g

b

g

b

g

p

p


MATLAB Simulation

The following function calculates the desired filter coefficients:

function v = relatflt(d1,d2)

%relatflt(d1,d2) len=7 FIR filter of response:

% 0  dB at f0 = 0.5

% d1 db at f1 = 0.25

% d2 db at f2 = 0.75

g1 = 10^(d1/20);

g2 = 10^(d2/20);

a  = (g1-g2)*0.37117514279802; % 3(2/8 - 1/2/(
b  = (g1+g2)*0.25 - 0.5;

c  = (g1-g2)*0.01762175220474; %  (2/8 - 1/2/(
d  = (g1+g2)*0.5;  

v  = [c b a  d  a b c];

Plot the filter response with linear compensation curves with gains [+0.4;–0.4], [+0.2;–0.2], [0;0], [–0.2;+0.2], [–0.4;+0.4]:

[image: image27.wmf]0

0.25

0.5

0.75

1

-0.5

-0.4

-0.3

-0.2

-0.1

0

0.1

0.2

0.3

0.4

0.5

Frequency

dB


and with decreasing curves: [+0.5;–0.4], [+0.3;–0.4], [0.1;–0.4], [–0.1;–0.4], [–0.3;–0.4]

[image: image28.wmf]0

0.25

0.5

0.75

1

-0.5

-0.4

-0.3

-0.2

-0.1

0

0.1

0.2

0.3

0.4

0.5

Frequency

dB

  

and with increasing curves: [–0.5;+0.4],
[–0.3;+0.4], [–0.1;+0.4], [+0.1;+0.4], [+0.3;+0.4]

[image: image29.wmf]0

0.25

0.5

0.75

1

-0.5

-0.4

-0.3

-0.2

-0.1

0

0.1

0.2

0.3

0.4

0.5

Frequency

dB


They agree completely what we wanted.

Implementation in C

The float d1 and d2 specify the desired filter gains in dB at 
[image: image30.wmf]f

1

1

4

=

 and 
[image: image31.wmf]f

2

3

4

=

. The filter coefficients are stored after calculation in the array Filt[Len], with Len = 7.
#define  Len 7

#define  C  ((Len-1)/2)

/*  from dB to ratio: */

float  g1 = pow( 10.0, d1 / 20.0);

float  g2 = pow( 10.0, d2 / 20.0);

float  Filt[Len];

/* 3(2/8 - 1/2/( = 0.37117514279802 */

/*  (2/8 - 1/2/( = 0.01762175220474 */

Filt[C-1] = Filt[C+1] = (g1-g2)*0.37117514279802;

Filt[C-2] = Filt[C+2] = (g1+g2)*0.25 - 0.5;

Filt[C-3] = Filt[C+3] = (g1-g2)*0.01762175220474;

Filt[ C ] =             (g1+g2)*0.5;

Extensions

The designed filters work well even beyond a ±6 dB compensation range. These large flatness errors, however, should not occur. They indicate serious hardware faults.

If the shape of the compensation curve needs to be more complex, we can easily add to the constraints a second pair of frequencies with specified gains. We should request the slope of the response curve at the innermost frequency points be equal to the slope of the secant going through the surrounding specified curve points. We need now a 15 tap filter of the form:


[image: image32.wmf]a

a

a

a

a

a

a

a

a

a

a

a

a

a

a

7

6

5

4

3

2

1

0

1

2

3

4

5

6

7

,

,

,

,

,

,

,

,

,

,

,

,

,

,

.

Let 
[image: image33.wmf]f

k

k

=

p

6

, (
[image: image34.wmf]k

=

1

5

,

K

) and the corresponding gain values 
[image: image35.wmf]g

k

, normalized at the center: 
[image: image36.wmf]g

3

1

=

. We have 5 equations requesting the correction gains:


[image: image37.wmf]g

a

a

i

k

k

i

i

=

+

F

H

G

I

K

J

=

=

å

0

1

7

2

6

1

5

cos

,

,

p

K


and another 3, requesting the correct slopes:


[image: image38.wmf]g

g

a

i

k

k

k

i

i

-

=

F

H

G

I

K

J

=

+

=

å

2

1

7

2

3

6

1

2

3

p

p

sin

,

,

,

.

The solution of the corresponding linear system of equations goes similarly as before:

[image: image39.wmf]a0

=

-

3

�

3

H

g1

-

2

g3

+

g5

L

+

H

5

g1

+

3

g2

-

4

g3

+

3

g4

+

5

g5

L

p

12

p

a1

=

J

3

H

g2

-

g4

L

+

�

3

H

g1

-

g5

L

N

H

-

2

+

p

L

4

p

a2

=

H

g1

-

2

g3

+

g5

L

J

-

3

�

3

+

4

p

N

12

p

a3

=

J

3

H

g2

-

g4

L

+

�

3

H

g1

-

g5

L

N

H

-

3

+

p

L

6

p

a4

=

H

g1

-

2

g3

+

g5

L

J

-

3

�

3

+

2

p

N

12

p

a5

=

72

g4

+

27

�

3

g5

+

36

g2

H

-

2

+

p

L

-

36

g4

p

-

5

�

3

g5

p

+

�

3

g1

H

-

27

+

5

p

L

72

p

a6

=

-

3

�

3

H

g1

-

2

g3

+

g5

L

+

H

g1

+

3

g2

-

8

g3

+

3

g4

+

g5

L

p

24

p

a7

=

-

3

J

3

g1

+

4

�

3

H

g2

-

g4

L

-

3

g5

N

+

J

g1

+

6

�

3

H

g2

-

g4

L

-

g5

N

p

24

�

3

p


 It is more complex, requiring a little more processor power for the run-time design, but the work is still manageable.

The MATLAB macro for the design:
function v = flt15(d1,d2,d4,d5)

%flt15(d1,d2,d4,d5) len=15 FIR filter, response:

% d1 db at f1 = 1/6

% d2 db at f2 = 2/6

% 0  dB at f0 = 3/6

% d4 db at f1 = 4/6

% d5 db at f2 = 5/6

g1 = 10^(d1/20); g4 = 10^(d4/20);

g2 = 10^(d2/20); g5 = 10^(d5/20);

a0 = 0.27883*(g1+g5) + 0.25000*(g2+g4) - 0.05767;

a1 = 0.15735*(g1-g5) + 0.27254*(g2-g4);

a2 = 0.19550*(g1+g5)                   - 0.39100;

a3 = 0.01301*(g1-g5) + 0.02254*(g2-g4);

a4 = 0.02883*(g1+g5)                   - 0.05767;

a5 =-0.08647*(g1-g5) + 0.18169*(g2-g4);

a6 =-0.02725*(g1+g5) + 0.12500*(g2+g4) - 0.19550;

a7 =-0.04486*(g1-g5) + 0.09085*(g2-g4);

v =[a7 a6 a5 a4 a3 a2 a1 a0 a1 a2 a3 a4 a5 a6 a7];

The following figure shows a few example response curves of filters designed by the above macro.

[image: image40.wmf]0

0.1667

0.3333

0.5

0.6667

0.8333

1

-0.5

-0.4

-0.3

-0.2

-0.1

0

0.1

0.2

0.3

0.4

0.5

Frequency

dB


Calibration Tables

Usually the input signal gets attenuated and/or amplified before the analog to digital conversion to assure maximum digital resolution. These attenuator-amplifier circuits affect each other to a different degree, dependent on the selected attenuation. In theory, we need a frequency response calibration table for each possible attenuation value.

If an instrument needs to be calibrated at several different attenuation levels over the whole frequency range, it will consume a lot of time and costs. Proper design reduces the number of necessary calibration runs.

In practice, without significant effort in the hardware design the frequency-charac​teristics change only at smaller attenuation values, because higher attenuations provide good de-coupling between otherwise interfering circuit parts.

It further reduces the number of necessary tables that the affects of de-coupled, cascaded attenuation devices are cumulative (additive when measured in dB).

Temperature compensation can be incorporated, too. The measured ambient or internal temperature represents another dimension for the family of tables. If the temperature dependency is smooth (linear or close to that) and does not change the shape of the frequency response curve, one extra table is enough for high precision compensation. The correction can be done by a temperature dependent multiplication factor (additive in dB). 

Literature

[1] A. D. Poularikas, Fomulas and Tables for Signal Processing, CRC Press – IEEE Press, 1999.

[2] N. Kalouptsidis, Signal Processing Systems: Theory and Design, John Wiley & Sons, 1997.

[3] J. G. Proakis, D. G. Manolakis: Digital Signal Processing: Principles, Algorithms and Applications, 3rd Edition, 1996 Prentice Hall.

[4] N. Fliege, Multirate Digital Signal Processing: Multirate Systems, filter Banks, Wavelets, John Wiley & Sons, 1995.

[5] W. H. Press, S. A. Teukolsky, W. T. Vetterling, B. P. Flannery: Numerical Recipes in C, Cambridge University Press, 1992.

[6] L. R. Rabiner, B. Gold: Theory and Application of the Digital Signal Processing, 1975 Prentice Hall.

[7] A. V. Oppenheim, R. W. Schafer: Digital Signal Processing, 1975 Prentice Hall.

Copyright  © 1997 Wavetek GmbH (Wireless Communications Division).
This document is company confidential.  It must not be copied in whole or in part or released to any third party without the prior written consent of Wavetek.

COMPANY CONFIDENTIAL




_995364082

_995364091

_995365914

_995469888

_995471606

_995471630

_995471234

_995467097

_995469080

_995469741

_995469783

_995469164

_995469024

_995373985

_995364155

_995365321

_995365765

_995365849

_995364454

_995364094

_995364096

_995364087

_995364089

_995364084

_995364066

_995364077

_995364080

_995364071

_995364059

_995364061

_995364057

